## Sheet1
CODE | Label | Description | Parameter | 单位 | 说明
1 | RED | Red Alarm, critical alarm, high level | alarm message |  | 字符串，支持多国语言
2 | ORANGE | Orange Alarm, serious alarm, medium level | alarm message |  | 字符串，支持多国语言
3 | YELLOW | Yellow Alarm, general alarm, low level | alarm message |  | 字符串，支持多国语言
101 | ECG | non-specific label for ECG. | wave
102 | I | Lead I - ECG specific label. | wave
103 | II | Lead II - ECG specific label | wave
104 | III | Lead III - ECG specific label | wave
105 | aVL | Augmented lead aVL | wave
106 | aVR | Augmented leads aVR | wave
107 | aVF | Augmented leads aVF | wave
108 | V | Chest lead V | wave
109 | V1 | Chest lead V1 | wave
110 | V2 | Chest lead V2 | wave
111 | V3 | Chest lead V3 | wave
112 | V4 | Chest lead V4 | wave
113 | V5 | Chest lead V5 | wave
114 | V6 | Chest lead V6 | wave
115 | HR | Heart Rate | numeric | bpm | -1为无效值
117 | ST1 | ST parameter - value ST1 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
118 | ST2 | ST parameter - value ST2 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
119 | ST3 | ST parameter - value ST3 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
120 | ST4 | ST parameter - value ST4 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
121 | ST5 | ST parameter - value ST5 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
122 | ST6 | ST parameter - value ST6 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
123 | ST7 | ST parameter - value ST7 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
124 | ST8 | ST parameter - value ST8 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
125 | ST9 | ST parameter - value ST9 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
126 | ST10 | ST parameter - value ST10 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
127 | ST11 | ST parameter - value ST11 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
128 | ST12 | ST parameter - value ST12 | numeric | mV | -1为无效值,=（收到的数值 – 256）/100
129 | PVCs | PVC sum per minute | numeric |  | -1为无效值,=（收到的数值 – 256）/100
201 | Resp | Respiration wave | wave
202 | RR | Respiration rate | numeric | rpm | -1为无效值
301 | Pleth | Pulse wave | wave
302 | SpO2 | Percent Oxyhemoglobin Saturation - Arterial | numeric | % | -1为无效值
303 | PR | Pusle Rate | numeric | bpm | -1为无效值
304 | PI | Perfusion Index | numeric | % | -1为无效值,=接到值/1000
401 | NIBP S | Non-invasive Blood Pressure Systolic | numeric | mmHg | -1为无效值
402 | NIBP D | Non-invasive Blood Pressure Diastolic | numeric | mmHg | -1为无效值
403 | NIBP M | Non-invasive Blood Pressure Mean | numeric | mmHg | -1为无效值
501 | T1 | Non-specific temperature 1 | numeric | ℃ | -1为无效值，=(收到值/10)℃
502 | T2 | Non-specific temperature 2 | numeric | ℃ | -1为无效值，=(收到值/10)℃
503 | TD | Delta Temperature between non-specific temperature labels | numeric | ℃ | -1为无效值，=(收到值/10)℃
601 | C.O. | Cardiac output | numeric | L/min | -1为无效值,=(收到值/100)
602 | C.I. | Cardiac Index | numeric |  | -1为无效值
603 | TB | Blood temperature - delivered by C.O. | numeric | ℃ | -1为无效值,=(收到值/100)
701 | ART | Specific Pressure Label - Arterial Blood Pressure | wave
702 | ART S | Specific Pressure Label - Arterial Blood Pressure Systolic | numeric | mmHg | -32768为无效值
703 | ART D | Specific Pressure Label - Arterial Blood Pressure Diastolic | numeric | mmHg | -32768为无效值
704 | ART M | Specific Pressure Label - Arterial Blood Pressure Mean | numeric | mmHg | -32768为无效值
705 | PA | Specific label - Pulmonary Artery Pressure | wave
706 | PA S | Specific label - Pulmonary Artery Pressure Systolic | numeric | mmHg | -32768为无效值
707 | PA D | Specific label - Pulmonary Artery Pressure Diastolic | numeric | mmHg | -32768为无效值
708 | PA M | Specific label - Pulmonary Artery Pressure Mean | numeric | mmHg | -32768为无效值
709 | CVP | Specific pressure label - Central Venuous Pressure | wave
710 | CVP M | Specific pressure label - Central Venuous Pressure Mean | numeric | mmHg | -32768为无效值
711 | LAP | Specific label - Left Atrial Pressure | wave
712 | LAP M | Specific label - Left Atrial Pressure Mean | numeric | mmHg | -32768为无效值
713 | RAP | Specific Pressure Label - Right Atrial Blood Pressure | wave
714 | RAP M | Specific Pressure Label - Right Atrial Blood Pressure Mean | numeric | mmHg | -32768为无效值
715 | ICP | Specific label - Intra Cranial Pressure | wave
716 | ICP M | Specific label - Intra Cranial Pressure Mean | numeric | mmHg | -32768为无效值
717 | P1 | non-specific label for Pressure 1 | wave
718 | P1 S | non-specific label for Pressure 1 Systolic | numeric | mmHg | -32768为无效值
719 | P1 D | non-specific label for Pressure 1 Diastolic | numeric | mmHg | -32768为无效值
720 | P1 M | non-specific label for Pressure 1 Mean | numeric | mmHg | -32768为无效值
721 | P2 | non-specific label for Pressure 2 | wave
722 | P2 S | non-specific label for Pressure 2 Systolic | numeric | mmHg | -32768为无效值
723 | P2 D | non-specific label for Pressure 2 Diastolic | numeric | mmHg | -32768为无效值
724 | P2 M | non-specific label for Pressure 2 Mean | numeric | mmHg | -32768为无效值
725 | P3 | non-specific label for Pressure 3 | wave
726 | P3 S | non-specific label for Pressure 3 Systolic | numeric | mmHg | -32768为无效值
727 | P3 D | non-specific label for Pressure 3 Diastolic | numeric | mmHg | -32768为无效值
728 | P3 M | non-specific label for Pressure 3 Mean | numeric | mmHg | -32768为无效值
729 | P4 | non-specific label for Pressure 4 | wave
730 | P4 S | non-specific label for Pressure 4 Systolic | numeric | mmHg | -32768为无效值
731 | P4 D | non-specific label for Pressure 4 Diastolic | numeric | mmHg | -32768为无效值
732 | P4 M | non-specific label for Pressure 4 Mean | numeric | mmHg | -32768为无效值
801 | CO2 | Carbon dioxide wave (airway) | wave
802 | EtCO2 | End tidal CO2 | numeric | mmHg | -1为无效值
803 | INSCO2 | Inspired Minimum CO2 | numeric | mmHg | -1为无效值
804 | FiCO2 | Inspired C2O | numeric | mmHg | -1为无效值
805 | O2 | O2 Wave | wave
806 | EtO2 | End-Tidal O2 | numeric | mmHg | -1为无效值
807 | FiO2 | Inspired O2 | numeric | mmHg | -1为无效值
808 | N2O | N2O Wave | wave
809 | EtN2O | End-Tidal N2O | numeric | % | -1为无效值
810 | FiN2O | Inspired N2O | numeric | % | -1为无效值
811 | AA | Anesthetic Agent Wave. Non-specific Label | wave
812 | EtAA | End-Tidal Anesthetic | numeric | % | -1为无效值，=接收值/10
813 | FiAA | Inspired Anesthetic Agent | numeric | % | -1为无效值，=接收值/10
814 | HAL | Halothane Wave | wave
815 | EtHAL | End-Tidal Halothane | numeric
816 | FiHAL | Inspired Halothane | numeric
817 | ENF | Enflurane Wave | wave
818 | EtENF | Inspired Enflurane | numeric
819 | FiENF | End-Tidal Enflurane | numeric
820 | ISO | Isoflurane Wave | wave
821 | EtISO | End-Tidal Isoflurane | numeric
822 | FiISO | Inspired Isoflurane | numeric
823 | SEV | Sevoflurane Wave | wave
824 | EtSEV | End-Tidal Sevoflurane | numeric
825 | FiSEV | Inspired Sevoflurane | numeric
826 | DES | Desflurane Wave | wave
827 | EtDES | End-Tidal Desflurane | numeric
828 | FiDES | Inspired Desflurane | numeric
829 | MAC | Minimum Alveolar Concentration | numeric | % | -1为无效值，=接收值/100
830 | AwRR | Air Way Respiration Rate | numeric | rpm | -1为无效值
901 | EEG | Inspired Sevoflurane | wave
902 | BIS | Sedation index | numeric | % | -1为无效值
903 | AI | Analgesic index | numeric | % | -1为无效值
1001 | reserved
1002 | reserved
1003 | reserved
8001 | PDF | report |  |  | PDF格式的报告，以base64格式封装
8002 | JPEG | report |  |  | JPEG格式的报告，以base64格式封装
8003 | XML | wave and other information |  |  | 波形数据或其它信息，以base64格式封装